## Slide 1
PostureSense
A Single-Accelerometer Wearable for Sleep Lying Position Detection
Hari Krishna Koneti · Northern Arizona University · Spring 2026
N=330 samples · 4 postures · Supine recall = 1.000 · 100% accuracy (controlled)
Arduino Uno R3 + ADXL345
~$30 BOM Cost
Open-Source Firmware
fₛ = 0.167 Hz
EE599 Wearable Computing · INF632 · Spring 2026

## Slide 2
Clinical and Engineering Motivation
Clinical Context
Positional OSA: AHI ratio supine/non-supine ≥2 [1,7]
Mechanism: gravity displaces tongue and soft palate posteriorly in supine posture, reducing airway patency [7]
Positional therapy (avoidance of supine sleep) is viable for a substantial patient subset
Overnight monitoring is essential to verify therapy adherence
Engineering Gap
No home-deployable device provides explicit lying-position labels
Polysomnography: laboratory-only, high cost, unsuitable for longitudinal monitoring
Commercial trackers infer sleep quality from cardiorespiratory proxies — none report supine duration directly
PostureSense: Explicit Posture Classification at ~$30 BOM Cost
Supine
Prone
Left Lateral
Right Lateral
N=330 samples · 4 postures · $30 BOM · supine recall = 1.000

## Slide 3
Review of Related Work
Alinia & Ghasemzadeh [2]
Chest-mounted single IMU · 89–95% accuracy across multiple subjects · Inter-subject placement variability is the dominant error source
Placement > Algorithm
Baan et al. [3]
Chest smartphone accelerometer · 90–96% accuracy · Sensor hardware quality is secondary to placement consistency
Cost-Feasible
Jeng et al. [4]
Wrist-mounted IMU · 88–93% in home-care setting · Wrist orientation varies with arm position, independent of trunk posture
Chest Preferred
Chen et al. [5]
Embedded threshold rules achieve accuracy comparable to trained classifiers for static posture detection on embedded hardware
Rule Sufficient
Fonseca et al. [8]
Controlled-to-free-living accuracy reduction: 5–15 pp · Controlled results represent an upper performance bound, not a deployment estimate
Controlled ≠ Overnight

## Slide 4
Research Question and Testable Hypotheses (α = 0.05)
RQ: Can a single tri-axial accelerometer, chest-mounted at fₛ = 0.167 Hz, discriminate four primary sleep lying positions exclusively from static gravitational acceleration components, achieving clinically meaningful classification performance?
H1 — Primary — Classifier Accuracy
H₀: Accuracy ≤ 0.90 or κ ≤ 0.85
H₁: Accuracy > 0.90 AND κ > 0.85
Evaluation: KNN (k=5), stratified 5-fold CV
Decision rule: reject H₀ at α = 0.05
Expected: gravitational class separability >> noise floor
H2 — Per-Class Sensitivity
H₀: Recall ≤ 0.90 for at least one class
H₁: Recall ≥ 0.90 for all four classes
Clinical priority: supine recall is primary metric
FN for supine = undetected OSA risk period
Decision rule: reject H₀ at α = 0.05
H3 — Embedded Rule Viability
H₀: Threshold classifier accuracy ≤ 0.85
H₁: Threshold classifier accuracy > 0.85
No stored model, no training data required
Rationale: Chen et al. [5]: threshold ≈ KNN
for static posture on embedded hardware
Decision rule: reject H₀ at α = 0.05

## Slide 5
Device Design and Firmware Configuration (sketch_mar24b.ino)
System Hardware
Microcontroller
Arduino Uno R3
Sensor
SparkFun ADXL345 breakout
Interface
I²C — SDA→A4, SCL→A5
Supply voltage
VCC → 3.3 V (3V3 rail)
Pull-up resistors
Onboard (SparkFun breakout)
Placement
Mid-sternum, medical adhesive
BOM cost
~$30 (component-level)
Dominant-Axis Mapping (T = 0.5 g):
Supine
a_y=+1g
Prone
a_y=-1g
Left Lat.
a_z=-1g
Right Lat.
a_z=+1g
Firmware Parameters (confirmed from sketch_mar24b.ino)
Firmware Call
Value
Implication
ADXL345 adxl =
ADXL345()
I²C mode
No CS pin → I²C default
adxl.setRangeSetting
(16)
±16 g
32 LSB/g scale factor
delay(6000)
6000 ms
fₛ = 0.167 Hz
Serial.print(
millis,x,y,z)
4 cols
Produces dataset columns
setActivityThreshold
(75)
4.69 g
ISR inactive; 4.5× max observed SMV
//attachInterrupt()
disabled
No effect on readings

## Slide 6
Experimental Methods and Dataset Characterisation
Data Collection Protocol
Study design:
Single-subject feasibility pilot (N=1)
Participant:
System developer (selection bias; see Limitations)
Surface:
Firm flat surface; no pillow (eliminates head-angle confound)
Protocol:
4 postures × ~3 min static hold each
Sensor placement:
Mid-sternum adhesive; not re-placed between postures
Firmware:
sketch_mar24b.ino — 0.167 Hz, raw counts via serial
Pre-processing:
Raw ADC / 32 → g-units | SMV = (ax²+ay²+az²)½
Labels:
Physics-based dominant-axis rule; no external ground truth
Analysis pipeline: Labels → Threshold (H3) → k sensitivity → 5-fold CV/KNN (H1,H2) → SMV ANOVA → Centroid
Chart: Sample Distribution (n=330, 32.9 min)
| Supine (N=57) | Prone (N=103) | Left Lat. (N=98) | Right Lat. (N=72)
Samples | 57 | 103 | 98 | 72
Class imbalance: prone/supine = 1.81×
Overall accuracy is potentially misleading; per-class recall is the primary evaluation metric.
Random seed = 42 · All code from scratch · No external ML libraries

## Slide 7
Feature Space Geometry and Class Separability
Pairwise Centroid Distances
Supine ↔ Prone
2.021 g
65×
Supine ↔ Left
1.328 g
43×
Supine ↔ Right
1.376 g
44×
Prone ↔ Left
1.342 g
43×
Left ↔ Right
1.772 g
57×
Physical Interpretation
ADXL345 noise floor (±16 g, 32 LSB/g): 0.031 g
Minimum centroid separation: 1.328 g
Separation ratio: 43× noise floor
Under controlled quasi-static conditions, correct classification is physically predicted by the sensor geometry.

## Slide 8
Result: Dominant-Axis Threshold Classifier — Hypothesis 3
Fig. 1 — Raw ADXL345 time series (a_x, a_y, a_z, SMV) from firmware serial output at fₛ = 0.167 Hz. Dashed lines mark posture-dependent dominant-axis transitions. Labels: physics-based rule.
H3 — Result
100.0%
330 / 330 correctly classified
0 ambiguous samples
Reject H₀ (H3) — Threshold Viable
7-branch rule · No learned parameters
T = 0.5 g = 16× above noise floor
Embedded: ATmega328P compatible
KNN adds no benefit (see k sensitivity)

## Slide 9
Result: KNN Classifier — Hypotheses 1 and 2 (k=5, Stratified 5-Fold CV)
Fig. 2 — Aggregated confusion matrix (5 CV folds). All 330 samples on the principal diagonal; FP = FN = 0 for all posture classes.
5-Fold Stratified CV
Overall accuracy:
1.0000 (100.0%)
Cohen’s κ:
1.0000
CV std deviation:
σ = 0.000 (no overfitting)
Reject H₀ (H1) — Accuracy + κ Both Satisfied
Per-Class Sensitivity (H2)
Supine
1.000
Clinical priority: FN=0
Prone
1.000
Left lateral
1.000
Right lateral
1.000
Reject H₀ (H2) — All Recall ≥ 0.90; Supine FN = 0

## Slide 10
Result: Complete Per-Class Performance Metrics
Fig. 3 — Precision, recall, specificity, and F1 per posture class. Red dashed: H1/H2 threshold (0.90). Derived from TP/FP/FN/TN across 5-fold CV aggregated confusion matrix.
Why Per-Class
Metrics Matter
Imbalance: prone/supine = 1.81×
Overall accuracy can mask poor per-class performance. A naive ‘predict prone always’ classifier achieves 31.2% overall while missing supine entirely.
Supine (N=57, 17.3%) is the critical clinical class.
All metrics = 1.000
for all 4 postures
FP = FN = 0 universally

## Slide 11
Result: Signal Magnitude Vector — One-Way ANOVA
Fig. 4 — SMV distributions per posture (one-way ANOVA, sum-of-squares decomposition, no external libraries). Dashed line: theoretical gravity baseline (1.000 g).
ANOVA Decomposition
F(3, 326):
184.52
p-value:
6.3 × 10⁻⁷⁰
SS between:
0.1420
SS within:
0.0836
Decision (α=0.05):
Reject H₀
Mean SMV per Posture
Supine
1.032 ± 0.007 g
Prone
0.995 ± 0.013 g
Left lat.
1.004 ± 0.006 g
Right lat.
0.967 ± 0.029 g
Physical basis [6]: supine → greater diaphragmatic excursion → SMV elevated. Effect size ΔSMV = 0.065 g = 2.1× noise floor.

## Slide 12
Discussion: Interpretation, Scope, and Experimental Design Limitations
Observations (What Data Show)
• 330 static samples from N=1 correctly classified by both approaches
• Centroid distances: 1.33–2.02 g (43–65× ADXL345 noise floor)
• SMV ANOVA: F(3,326)=184.52, p=6.3×10⁻⁷⁰ — posture-dependent respiratory motion confirmed
• Effect size: ΔSMV = 0.065 g = 2.1× noise floor (significant; modest practical magnitude)
Physical Inference (Mechanism)
• Gravitational geometry predicts classification: gravity uniquely constrains each posture to a dominant accelerometer axis
• Threshold rule exploits structural physics, not statistical regularities
• 43× noise margin renders class overlap geometrically implausible under quasi-static conditions
• SMV encodes respiratory excursion intensity in addition to posture identity
Scope Boundary (Critical Caveat)
• Results scoped to N=1, controlled static conditions only
• 100.0% accuracy is an upper performance bound, not a deployment specification
• Expected free-living performance: 86–94% [2,8]
• Generalisation requires independent multi-subject overnight validation (N ≥ 5)
Experimental Design Limitations
1. Activity interrupt threshold (setActivityThreshold(75) = 4.69 g) is 4.5× the maximum observed SMV — ISR never activated; replaced by delay(6000) polling.
2. k sensitivity (k = 1–11) yielded invariant accuracy — KNN provides no classification benefit over the threshold rule for this geometrically separable dataset.

## Slide 13
Contextualisation: Comparison to Commercial Sleep Monitoring Devices
Feature
PostureSense (This Work)
Withings Sleep Analyzer
Garmin Vivosmart 5
ResMed S+
Company / Year
NAU Research, 2026
France, est. 1998
Switzerland, est. 1989
USA, est. 1989
Sensing modality
Chest accelerometer
Pressure mat (BCG)
Wrist accel + HR
Bedside radar
Explicit posture output
✓ Yes — 4 labelled classes
✗ None
✗ None
✗ None
Component cost
~$30
~$130
~$150
~$150
Firmware / source
Open-source (GitHub)
Proprietary
Proprietary
Proprietary
Data storage
Local CSV
Cloud (opt-in)
Cloud (Garmin Connect)
Cloud (required)
OSA posture flag
Supine recall = 1.000
Indirect via HR/SpO₂
Via sleep score
Via breathing rate
Multi-subject validated
No (pilot only)
Yes
Yes
Yes
PostureSense is the only evaluated system providing explicit, labelled lying-position output — directly addressing the positional OSA monitoring requirement.

## Slide 14
Limitations and Threats to Validity
1 Single Subject (N=1) — Selection Bias
Developer-as-participant: optimal placement and deliberate static holds inflate controlled accuracy
Expected free-living accuracy degradation: 5–12 pp [2] → estimated overnight range: 86–94%
Ground truth derived from sensor data — no independent reference (video/PSG)
2 Controlled Static Protocol — No Transitions
Dataset contains no posture transitions — primary anticipated source of overnight error
At fₛ = 0.167 Hz: each 5–15 s transition generates 1–2 misclassified samples
Instantaneous feature set violates quasi-static assumption during transitions
3 Sensor Configuration and Latency
±16 g range → 32 LSB/g (sufficient for posture; ±2 g gives 8× finer resolution at no cost)
fₛ = 0.167 Hz → 6-second detection latency (acceptable for posture; inadequate for respiratory monitoring)
±2 g setting recommended for future iterations to improve SMV respiratory sensitivity
4 Overfitting and Generalisability
Overfitting tested via 5-fold CV: σ = 0.000 — no overfitting detected in this dataset
Zero CV variance expected: 43× geometric separation renders fold composition irrelevant
Clinical generalisation requires independent multi-subject overnight evaluation (N ≥ 5)
5 Calibration: No external angular reference. A 5° placement deviation shifts baseline a_y by sin(5°) ≈ 0.09 g, reducing threshold margin from 0.50 g to 0.41 g.

## Slide 15
Conclusions
Hypothesis Outcomes
H3 threshold accuracy
100.0%
Reject H₀
H1 KNN accuracy (5-fold CV)
100.0%
Reject H₀
H1 Cohen’s κ
1.0000
Reject H₀
H1/H2 CV std deviation
σ = 0.000 (no overfitting)
Verified
H2 supine recall (clinical primary)
1.000 — FN = 0
Reject H₀
SMV ANOVA F(3,326)
184.52, p=6.3×10⁻⁷⁰
Sig. (α=0.05)
Min. centroid separation
1.33 g = 43× noise floor
Physical basis
BOM cost
~$30
Feasible
H₀ Rejected for H1, H2, H3 — All Three Hypotheses Supported
The One Takeaway
A single $8 accelerometer mounted on the chest detects every supine sleep period with zero false negatives under controlled conditions.
The technology barrier is not the hardware. It is the 5–12 pp gap between controlled laboratory performance and real overnight reliability across diverse patients.
Closing that gap is the essential next step.

## Slide 16
References (IEEE Format)
[1]
V. K. Kapur et al., Clinical practice guideline for diagnostic testing for adult OSA. J. Clin. Sleep Med. 13(3):479-504, 2017.
[2] †
P. Alinia and H. Ghasemzadeh, Pervasive lying posture tracking. Sensors 20(20), 2020.
[3]
J. Baan et al., Lying position classification using accelerometry for positional OSA. Sensors 19(1), 2019.
[4]
P.-Y. Jeng et al., A wrist sensor sleep posture monitoring system. Sensors 21(1), 2021.
[5]
Y. Chen et al., Body position detection and posture reconstruction based on wearable IMUs. J. Med. Biol. Eng. 36(3), 2016.
[6] †
J. Razjouyan et al., Improving sleep quality assessment using wearable sensors. J. Clin. Sleep Med. 13(11):1301-1310, 2017.
[7]
D. A. Pevernagie et al., On the rise and fall of the positional OSA syndrome. Sleep Med. Rev. 14(6):371-375, 2010.
[8]
P. Fonseca et al., Automatic sleep staging using HR variability and body movement. Physiol. Meas. 36(11), 2015.
† Describes accelerometer use in sleep research context — satisfies the ≥2 research-context citation requirement.

## Slide 17
Thank You
Questions?
Supine recall:
1.000 — Zero false negatives; no OSA risk period missed
Hardware:
Arduino Uno R3 + SparkFun ADXL345 (~$8 sensor, $30 BOM)
Code & data:
github.com/HariKrishnaKonetii/Wearable-Computing